1. The Medical Devices Directive (MDD)

Definition and Purpose:

The MDD (Directive 93/42/EEC) provides the regulatory framework for medical devices within the EU.

Purpose:

1. Harmonize regulations across Member States to ensure the free movement of medical devices within the EU.

2. Protect public health by setting minimum safety and performance standards for medical devices.

Scope:

- Covers instruments, appliances, software, and other devices intended by manufacturers for medical purposes, such as diagnosis, prevention, treatment, or alleviation of diseases.

- Includes devices used for monitoring and controlling conception.

Usage:

- Devices that meet the directive's requirements are awarded a CE marking, which allows them to be marketed across the EU without additional national barriers.

- The directive applies to both physical devices and stand-alone software, provided they serve a specific medical purpose.

Core Interpretation of Article 1(1) and 1(2)(a) of the MDD

General Scope:

Article 1(1) defines the directive's scope, stating that it applies to medical devices intended by the manufacturer for specific medical purposes, such as diagnosis, prevention, monitoring, or treatment of diseases.

- This scope covers all types of medical devices, including physical instruments, tools, and stand-alone software.

Specific Relevance to Software:

Article 1(2)(a) provides a definition of what constitutes a "medical device." It states that any instrument, apparatus, or software intended for medical purposes qualifies as a medical device.

1. Uses patient-specific data: The software must process data tailored to individual patients.

2. Serves a medical purpose: For example, detecting contraindications, drug interactions, or excessive dosages.

3. Does not act directly in or on the human body: Unlike traditional medical devices like implants or surgical tools, software may operate indirectly by providing diagnostic or decision-support information.

2. The Medical Devices Regulation (MDR)

Adopted in 2017, the MDR replaced the MDD in 2020.

1. Art.2: Expanded definitions to include software used for prediction and prognosis in healthcare.

2. MDR states in its Recital 19 that ‘software for general purposes, even when used in a health care setting, or software intended for lifestyle and well-being purposes is not a medical device’.

3. SNITEM Case: Facts, Issues, and Arguments

Facts:

Parties:

SNITEM: A French medical technology association.

Philips France: A manufacturer of the Intellispace Critical Care and Anesthesia (ICCA) software.

French government and its Haute Autorité de Santé (High Health Authority), which required additional certification for CE-marked medical software.

Dispute:

ICCA is a prescription assistance software that analyzes patient data to detect contraindications, interactions, and dosages.

While the software had a CE marking, the French decree required it to undergo additional national certification.

SNITEM and Philips France argued this breached the MDD, which prohibits Member States from imposing barriers to CE-marked devices.

Legal Question:

Does software that does not act directly on the human body but supports medical decisions qualify as a medical device under Article 1(2)(a) of the MDD?

Advocate General’s (AG) Opinion

The Advocate General, Campos Sánchez-Bordona, issued a non-binding opinion that guided the court’s decision.

Key Points:

1. Definition of a Medical Device:

- Must have a medical purpose, such as use in humans for the purposes of diagnosis, prevention, monitoring, treatment, or alleviation of disease.

- In particular by means of calculation, quantification or comparison of the recorded data against certain references in order to provide information about a particular patient.

- Includes software for the ‘prediction’ and ‘prognosis’ of disease in the medical device definition in Article 2 MDR

2. CE Marking:

- Implies conformity with EU standards, allowing free movement within the EU.

- Additional national certification undermines the directive’s harmonization goal.

Analysis of ICCA Software:

- The ICCA software aids in prescribing drugs by analyzing patient-specific data, aligning with the directive’s purpose.

- Goes beyond simple data storage, providing actionable medical information.

CJEU Judgment

The CJEU followed the AG’s reasoning and ruled that:

1. Intended Purpose:

Software qualifies as a medical device if the manufacturer declares a medical purpose, such as diagnosis or treatment, even if it does not act directly on the human body.

2. Free Movement of Goods:

CE-marked devices must circulate freely across the EU without additional national barriers unless public health risks are established.

Application to ICCA:

The ICCA software met the criteria as it processed patient-specific data to assist in diagnosing and prescribing, fulfilling a therapeutic purpose.

Reasoning

1. Manufacturer's Intent:

The intended use of the device, as defined by the manufacturer, determines whether it qualifies as a medical device.

2. Functionality:

Stand-alone software that processes data to provide medical insights is distinct from software performing administrative tasks like storage or communication.

3. Regulatory Harmonization:

Additional certification undermines the harmonization goals of the MDD, as CE marking already ensures compliance with EU standards.

Supporting Case Law:

1. Brain Products Case (C-219/11):

Established that software must serve a medical purpose to qualify as a medical device.

2. Oliver Medical Case (C-547/13):

CE marking would imply that the product is for medical use

3. James Elliott Construction Case (C-613/14)

The CJEU ruled that the CE marking creates a presumption of conformity with the essential requirements of the applicable EU directive

Implications of the Decision

1. For Manufacturers:

- Software developers must clearly define the medical purpose of their products to qualify as medical devices.

- CE marking ensures access to the EU market without additional national requirements.

3. For Regulators:

- National authorities must adhere to EU harmonization principles and cannot impose unnecessary barriers.

- Regulators must balance safety concerns with innovation and market accessibility.

3. For Digital Health:

- The decision sets a precedent for regulating AI and machine learning in healthcare.

- Encourages innovation by providing legal clarity for stand-alone medical software.

4. National Impact: French Conseil d’État Decision

Following the CJEU ruling, France’s Conseil d’État annulled parts of Decree No. 2014-1359 that imposed additional certification requirements on CE-marked software.

This decision aligned French law with EU principles, ensuring the free movement of medical devices.

4. Broader Impact

1. Legal Clarity:

The ruling strengthens the regulatory framework for digital health tools, aligning with the MDR’s expanded definitions.

2. Emerging Technologies:

AI-driven software and big data applications in healthcare will benefit from the clarified criteria for classification as medical devices.