### 1. Legal Sources, Software, and Use in the Medical Industry

#### 1. Medical Device Directive (MDD)

The MDD, effective until May 2021, had limited guidance for software, often requiring physical integration with devices to qualify as an MD. Article 1(2) defined MD purposes, including:

- Diagnosis, prevention, monitoring, treatment, or alleviation of disease.

- Diagnosis, monitoring, treatment, alleviation, or compensation for an injury or handicap.

- Investigation, replacement, or modification of anatomy or physiological processes.

- Control of conception.

Software had to fulfill one of these purposes to qualify as an MD, excluding general-purpose tools unless explicitly customised for medical use. Fragmentation in national implementation under the MDD necessitated the MDR’s harmonised framework.

#### 2. Medical Device Regulation (MDR)

#### Standalone software refers to software not embedded in or dependent on a physical medical device for functionality.

The MDR replaced the MDD in May 2021, directly enforceable across Member States, and expanded the scope of MD regulation. Article 2(1) defines MD purposes with slight updates:

- Diagnosis, prevention, monitoring, prediction, prognosis, treatment, or alleviation of disease.

- Diagnosis, monitoring, treatment, alleviation, or compensation for an injury or disability.

- Investigation, replacement, or modification of anatomy or physiological or pathological processes or states.

- Providing information via in vitro examination of specimens derived from the human body.

The MDR’s broader scope explicitly includes standalone software, provided it fulfills these purposes, even without physical interaction with the body.

#### 3. Guidance Documents

- MEDDEV 2.1/6 (MDD Guidance): MEDDEV 2.1/6 distinguishes between software that stores, archives, communicates or searches in data and software that does other actions. There is no basis for this in the Regulation, and the points above support the idea, since purely storing, archiving, communicating or searching does not fulfil any of the purposes.

- MDCG 2019-11 (MDR Guidance):

(1) The first part relates to whether the software can be defined as software in the guidance. Related to the MDR, there is no stringent definition of this. This is in itself tautological.

(2) The second part relates to whether the software is standalone or not. If it is not, it can either be part of a MD or not covered by the MDD.

(3) The third part defines that if the software merely acts on data from storage, archives, communication or simple search, it is not a MD.

(4) The fourth part postulates that if the actions of the software are not for the benefit of individual patients, it is not a MD. Note that neither the MDD nor the MDR has the benefit of patients among their explicit main purposes.46

(5) The fifth part says that if the software does not fit a purpose contained in Article 1(2)(a)of the MDD, it is not a MD, and it is if it does. But even if it is not a MD, the software can still be considered an accessory, leading to the same requirements as that of a MD.

(6) The sixth part says that in order to be an accessory,the software must fulfil Article 1(2)(b), which means that it must be intended specifically to be used with an MD by the manufacturer. Otherwise, it is not a MD or an accessory at all.

Classification

#### - Rule 11 in Annex VIII assumes at first that all software that is also a MD is class I.

#### - The exception to this is software that is used to take decisions on diagnosis or has therapeutic purposes. These are instead considered to be class IIa.

#### - The exception to the exception would be if the software can cause serious deterioration of the health of a person, which makes it class IIb

#### - It is capable of causing death or irreversible deterioration to the health of a person, in which case it is class III.

#### - If the software monitors physiological processes, it is class IIa, but if it can cause immediate harm to the patient during its use, it is to be considered class IIb.

#### 4. Case Law: C-329/16 SNITEM

This pivotal CJEU case addressed whether software supporting medical decisions qualifies as an MD. Key rulings:

- Standalone software aiding medical decision-making qualifies as an MD if it fulfils medical purposes under Article 1(2).

- Physical interaction with the body is unnecessary; classification hinges on the intended purpose of the manufacturer.

### 2. The Intention of the Manufacturer

#### 1. Challenges in Determining Intention

The MDR emphasizes intended purpose as a gatekeeper for MD classification. However, manufacturers can circumvent regulation by omitting medical claims in public-facing materials while designing software for medical uses.

#### 2. Framework for Determining Intention

The journal proposes a two-pronged framework to assess direct and indirect intention:

##### Direct Intention

- Marketing Materials: Explicit claims about fulfilling Article 2(1) purposes establish direct intention.

- Internal Documentation: Statements in technical or internal documents accessible to regulators affirm medical intent.

- Informal Sources: Metadata, tags, or unofficial statements indicating medical use.

##### Indirect Intention

- Data-Gathering Practices: Collecting biometrical or health-related data relevant to medical purposes.

- Data Analysis: Processing personal data to mimic medical diagnostics or treatments.

- Functionality: Replicating or replacing MDs (e.g., prosthetic controllers, diagnostic apps).

Key Insight: If either type of intention is evident, further evaluation determines compliance with Article 2(1) purposes.

### 3. Requirements for Software to Be a Medical Device

#### 1. Overarching Concepts

Software qualifies as an MD if it:

- Fulfills at least one Article 2(1) purpose. These include diagnosis, monitoring, or compensation for injury/disability.

- Demonstrates intended medical use. This can be direct (explicit) or indirect (implied through functionality).

- Is not generic software. Only modified software designed for specific medical purposes qualifies.

- Is human-centric. As clarified in SNITEM, software must benefit or act on humans, even without direct physical interaction.

#### 2. Considerations for Special Cases

- Excluded Software: Tools solely storing, archiving, or communicating data without medical analysis are excluded.

- Accessory Software: Software supporting MDs (e.g., controlling surgical robots) is classified as an accessory.

- Risk-Based Classification: MDR Annex VIII divides devices (including software) into classes I-III based on risk, with stricter rules for higher-risk categories (e.g., diagnostic tools that could cause harm).

### 4. Proposed Reforms

#### 1. Broadened Interpretation of Intention

- Expand the definition of intention to include indirect uses and inferred purposes, preventing circumvention.

- Ensure accessory software enabling MD functionalities is regulated appropriately.

#### 2. Enhanced Surveillance and Enforcement

- Regulators should proactively investigate software markets, focusing on indirect functionalities.

- Strengthen oversight of software distributed on platforms like app stores.

#### 3. Comprehensive Risk Mitigation

- Mandate robust cybersecurity and data privacy protections for SaMD.

- Require periodic updates and recertifications to maintain compliance.